## Bioanalytical method
Bioanalytical method description
Study name
Study code
Active ingredient/s
Instrument name
Freezer type and code
Solvents(add rows if required) | Description
Parameters | Description
Mobile Phase
Analytical column/code
Guard column
Autosampler tempreature
Flow rate
Column tempreture
Injection volume
Total run time
Processing parameters
Analyte
standard code
RT window
Smooth Width
Noise percent
Base Sub window
Peak splitting factor
Internal standard
Internal standard code
RT window
Smooth Width
Noise percent
Base Sub window
Peak splitting factor
Chromatogramic features
Compound
Retention time
IS
Retention time
MS conditions
Analyte(s) | Drug(s) | IS | -28 | -28 | -10
Q1 (m/z) |  |  | -24 | -24 | -19
Q3 (m/z)
DP (v)
EP (v)
CE (v)
CXP (v)
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
Lab Technical Manager
Date
IPC Specialist
Date
SOP/LB/07/F01/V1.0 Eff. Date: 22/4/2018

## Linearity
Linearity Report
active ingredient/s | [24 empty] | Mean of Accurcey% | standard deviation | CV% | Accepted/Rejected
Calibration no. | 1 |  |  |  | 2 |  |  |  | 4 |  |  |  | 5 |  |  |  | 8 |  |  |  | 9
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser type and code | Microlit - Zi\DI\MI\02 (10 mL) |  |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  |  | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge type and code | Eppendorf - Zi/CE/EP/01 |  |  |  | Eppendorf - Zi/CE/EP/01 |  |  |  | Eppendorf - Zi/CE/EP/01 |  |  |  | Eppendorf - Zi/CE/EP/01 |  |  |  | Eppendorf - Zi/CE/EP/01 |  |  |  | Eppendorf - Zi/CE/EP/01
Concentrator type and code | N/A |  |  |  | N/A |  |  |  | N/A |  |  |  | N/A |  |  |  | N/A |  |  |  | N/A
Balance type and code | Ohous - Zi/BA/OH/01 |  |  |  | Ohous - Zi/BA/OH/01 |  |  |  | Ohous - Zi/BA/OH/01 |  |  |  | Ohous - Zi/BA/OH/01 |  |  |  | Ohous - Zi/BA/OH/01 |  |  |  | Ohous - Zi/BA/OH/01
Date | 2019-07-04 |  |  |  | 2019-07-05 |  |  |  | 2019-07-05 |  |  |  | 2019-07-05 |  |  |  | 2019-07-08 |  |  |  | 2019-07-09
Batch start time | 23:57 |  |  |  | 1:39 |  |  |  | 14:34 |  |  |  | 16:37 |  |  |  | 10:25 |  |  |  | 17:55
Batch end time | 1:06 |  |  |  | 3:37 |  |  |  | 15:49 |  |  |  | 17:36 |  |  |  | 11:24 |  |  |  | 18:54
Batch name
Expected value (Calibration levels) (in µg/ml) | Peak area ratio | Found conc (in ng/ml) | Accuracy (%) | Accepted/Rejected | Peak area ratio | Found conc (in ng/ml) | Accuracy (%) | Accepted/Rejected | Peak area ratio | Found conc (in ng/ml) | Accuracy (%) | Accepted/Rejected | Peak area ratio | Found conc (in ng/ml) | Accuracy (%) | Accepted/Rejected | Peak area ratio | Found conc (in ng/ml) | Accuracy (%) | Accepted/Rejected | Peak area ratio | Found conc (in ng/ml) | Accuracy (%) | Accepted/Rejected |  | STD | CV% | Accepted/Rejected
0.5 | 0.060247 | 0.525 | 105.011 | Accepted | 0.061294 | 0.517 | 103.325 | Accepted | 0.044345 | 0.499 | 99.773 | Accepted | 0.066781 | 0.527 | 105.377 | Accepted | 0.059904 | 0.521 | 104.277 | Accepted | 0.06063 | 0.52 | 104.09 | Accepted | 103.6422 | 2.0277 | 1.9564 | Accepted
2 | 0.29054 | 1.807 | 90.357 | Accepted | 0.28081 | 1.842 | 92.12 | Accepted | 0.22712 | 2.007 | 100.348 | Accepted | 0.28449 | 1.762 | 88.114 | Accepted | 0.28138 | 1.826 | 91.303 | Accepted | 0.26863 | 1.802 | 90.108 | Accepted | 92.0583 | 4.279 | 4.6481 | Accepted
3 | 0.44289 | 2.655 | 88.51 | Accepted | 0.43982 | 2.803 | 93.428 | Accepted | 0.33573 | 2.903 | 96.771 | Accepted | 0.4284 | 2.579 | 85.963 | Accepted | 0.41872 | 2.635 | 87.837 | Accepted | 0.40347 | 2.633 | 87.769 | Accepted | 90.0463 | 4.1415 | 4.5992 | Accepted
5 | 0.77204 | 4.488 | 89.755 | Accepted | 0.80418 | 5.003 | 100.069 | Accepted | 0.65655 | 5.55 | 111.006 | Accepted | 0.81364 | 4.765 | 95.298 | Accepted | 0.78497 | 4.793 | 95.852 | Accepted | 0.77888 | 4.947 | 98.93 | Accepted | 98.485 | 7.1103 | 7.2197 | Accepted
10 | 1.4264 | 8.131 | 81.309 | Rejected | 1.4332 | 8.802 | 88.023 | Accepted | 1.0851 | 9.087 | 90.865 | Accepted | 1.7904 | 10.307 | 103.075 | Accepted | 1.7229 | 10.318 | 103.176 | Accepted | 1.7209 | 10.752 | 107.516 | Accepted | 95.6607 | 10.3856 | 10.8567 | Accepted
30 | 5.8724 | 32.882 | 109.608 | Accepted | 5.6555 | 34.304 | 114.347 | Accepted | 3.7242 | 30.861 | 102.872 | Accepted | 5.4343 | 30.984 | 103.281 | Accepted | 5.0609 | 29.981 | 99.938 | Accepted | 5.0513 | 31.274 | 104.246 | Accepted | 105.7153 | 5.2765 | 4.9912 | Accepted
45 | 8.8973 | 49.722 | 110.494 | Accepted | 7.8549 | 47.588 | 105.751 | Accepted | 5.5738 | 46.123 | 102.497 | Accepted | 8.4116 | 47.879 | 106.398 | Accepted | 7.963 | 47.077 | 104.615 | Accepted | 7.8147 | 48.303 | 107.34 | Accepted | 106.1825 | 2.6908 | 2.5341 | Accepted
50 | 9.5098 | 53.132 | 106.265 | Accepted | 8.4974 | 51.468 | 102.937 | Accepted | 5.7933 | 47.934 | 95.868 | Accepted | 9.8863 | 56.247 | 112.495 | Accepted | 9.5629 | 56.502 | 113.003 | Accepted | 1.7549 | 10.961 | 21.922 | Rejected | 92.0817 | 34.9575 | 37.9636 | Rejected
Correlation equation | 0.179625 * x + -0.0340656 |  |  |  | 0.165571 * x + -0.0242436 |  |  |  | r = 0.997830, r^2 = 0.995665 |  |  |  | 0.176228 * x + -0.0260711 |  |  |  | 0.169757 * x + -0.0286049 |  |  |  | 0.162278 * x + -0.0238277
r | r = 0.993847, r^2 = 0.987732 |  |  |  | r = 0.995554, r^2 = 0.991127 |  |  |  | 0.121195 * x + -0.0161153 |  |  |  | r = 0.994690, r^2 = 0.989408 |  |  |  | r = 0.995970, r^2 = 0.991956 |  |  |  | r = 0.995834, r^2 = 0.991685
(Linear, weighed 1/X2,….) | 1/X2 |  |  |  | 1/X2 |  |  |  | 1/X2 |  |  |  | 1/X2 |  |  |  | 1/X2 |  |  |  | 1/X2
Points excluded | 1 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  |  |  | 1
% of accepted calibrators | 87.5% |  |  |  | 100% |  |  |  | 100% |  |  |  | 100% |  |  |  | 100% |  |  |  | 87.5%
Reason for exclusion | Dosen`t fullfill acceptance cariteria
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F02/V1.0 Eff. Date: 22/4/2018
Blank acceptance criteria

## Accuracy
Accuracy report
active ingredient/s |  | Meropemene
Inter day Accuracy
Calibration no. | 1
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used and Code | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A
Batch name | 1st day validation ( Cal-1 )
Date | 2019-07-04
Day | One
Level | LLOQ |  |  | Low QC |  |  | Medium QC |  |  | High QC
Accepted range | ± 20 % |  |  | ± 15 % |  |  | ± 15 % |  |  | ± 15 %
Nominal value | 0.5 |  |  | 1.5 |  |  | 15 |  |  | 40
SN# | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail
1 | [0.608] | [121.594] | Fail | 1.344 | 89.61 | Pass | 13.802 | 92.015 | Pass | 37.975 | 94.936 | Pass
2 | 0.582 | 116.461 | Pass | 1.28 | 85.311 | Pass | 13.703 | 91.35 | Pass | 40.127 | 100.316 | Pass
3 | 0.559 | 111.794 | Pass | [1.195] | 79.665 | Fail | 13.412 | 89.41 | Pass | 39.848 | 99.621 | Pass
4 | [0.623] | [124.668] | Fail | [1.222] | 81.464 | Fail | 13.163 | 87.755 | Pass | 39.193 | 97.983 | Pass
5 | 0.584 | 116.819 | Pass | 1.294 | 86.275 | Pass | 14.413 | 96.088 | Pass | 41.657 | 104.144 | Pass
6 | 0.585 | 117.098 | Pass | 1.299 | 86.613 | Pass | 14.158 | 94.389 | Pass | 42.286 | 105.715 | Pass
Mean | 0.5775 |  |  | 1.3043 |  |  | 13.7752 |  |  | 40.181
CV% | 2.1465 |  |  | 2.1233 |  |  | 3.3519 |  |  | 3.9454
Mean Accuracy% | 115.5 |  |  | 86.95 |  |  | 91.8344 |  |  | 100.4525
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F03/V1.0 Eff. Date: 22/4/2018
active ingredient/s |  | Meropemene
Inter day Accuracy
Calibration no. | 4
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used and Code | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A
Batch name | 2nd day validation ( Cal -4 )
Date | 2019-07-05
Day | Two
Level | LLOQ |  |  | Low QC |  |  | Medium QC |  |  | High QC
Accepted range | ± 20 % |  |  | ± 15 % |  |  | ± 15 % |  |  | ± 15 %
Nominal value | 0.5 |  |  | 1.5 |  |  | 15 |  |  | 40
SN# | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail
1 | 0.523 | 104.621 | Pass | 1.294 | 86.27 | Pass | 13.968 | 93.122 | Pass | 41.65 | 104.126 | Pass
2 | 0.424 | 84.744 | Pass | [1.185] | [79.028] | Fail | 14.405 | 96.035 | Pass | 40.158 | 100.396 | Pass
3 | 0.557 | 111.404 | Pass | 1.28 | 85.365 | Pass | [12.483] | [83.222] | Fail | 39.026 | 97.566 | Pass
4 | 0.582 | 116.436 | Pass | [1.218] | [81.222] | Fail | 14.15 | 94.337 | Pass | 40.548 | 101.369 | Pass
5 | 0.574 | 114.863 | Pass | 1.277 | 85.101 | Pass | 14.604 | 97.363 | Pass | 40.85 | 102.125 | Pass
6 | 0.57 | 113.996 | Pass | 1.284 | 85.588 | Pass | 15.165 | 101.102 | Pass | 42.935 | 107.337 | Pass
Mean | 0.53833 |  |  | 1.2837 |  |  | 14.4584 |  |  | 40.8612
CV% | 11.095 |  |  | 0.57726 |  |  | 3.205 |  |  | 3.2621
Mean Accuracy% | 107.6667 |  |  | 85.5833 |  |  | 96.3893 |  |  | 102.1529
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F04/V1.0 Eff. Date: 22/4/2018
active ingredient/s |  | Meropemene
Inter day Accuracy
Calibration no. | 1
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used and Code | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A
Batch name | 3rd day validation ( Cal-7 )
Date | 2019-07-08
Day | Three
Level | LLOQ |  |  | Low QC |  |  | Medium QC |  |  | High QC
Accepted range | ± 20 % |  |  | ± 15 % |  |  | ± 15 % |  |  | ± 15 %
Nominal value | 0.5 |  |  | 1.5 |  |  | 15 |  |  | 40
SN# | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail | Actual value | A % | Pass/Fail
1 | 0.567 | 113.446 | Pass | 1.318 | 87.867 | Pass | 12.852 | 85.682 | Pass | 35.258 | 88.146 | Pass
2 | 0.527 | 105.368 | Pass | 1.287 | 85.793 | Pass | 13.145 | 87.633 | Pass | 34.621 | 86.552 | Pass
3 | 0.545 | 108.958 | Pass | 1.305 | 87.032 | Pass | [12.57] | [83.797] | Fail | 36.826 | 92.065 | Pass
4 | 0.532 | 106.3 | Pass | [1.234] | [82.265] | Fail | 13.282 | 88.546 | Pass | 38.897 | 97.243 | Pass
5 | 0.519 | 103.833 | Pass | 1.313 | 87.558 | Pass | 13.297 | 88.644 | Pass | 37.733 | 94.332 | Pass
6 | 0.527 | 105.322 | Pass | [1.242] | [82.796] | Fail | 13.269 | 88.458 | Pass | 38.211 | 95.528 | Pass
Mean | 0.53617 |  |  | 1.3057 |  |  | 13.169 |  |  | 36.9243
CV% | 3.2399 |  |  | 1.0414 |  |  | 1.4217 |  |  | 4.5789
Mean Accuracy% | 107.2333 |  |  | 87.05 |  |  | 87.7933 |  |  | 92.3108
Inter day Accuracy for 3 days
Level | LLOQ |  |  | Low QC |  |  | Medium QC |  |  | High QC
Conc | 0.5 |  |  | 1.5 |  |  | 15 |  |  | 40
Mean Concentration | 0.54731 |  |  | 1.2979 |  |  | 13.7993 |  |  | 39.3222
Mean Accuracy R%
SD | 0.040564 | 4.9978 | #DIV/0! | 0.019607 | 2.713 | #DIV/0! | 0.64431 | 4.6696 | #DIV/0! | 2.2865 | 4.2323 | #DIV/0!
CV% | 7.4115 |  |  | 1.5107 |  |  | 4.6692 |  |  | 5.8148
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F05/V1.0 Eff. Date: 22/4/2018

## Selectivity
Selectivity
active ingredient/s | Meropeneme
Calibration no. | 2
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL)
Concentrator code | Eppendorf - Zi/CE/EP/01
Centrifuge code | N/A
Batch name | 1st day validation ( Cal -2 )
Date | 2019-07-05
SN | Area of drug at LLOQ | Area of interfering peak of drug | % | Pass/Fail | area of Internal standard | Area of interfering peak of Internal | % | Pass/Fail
blank plasma batch 1 | 664.827 | 15.586 | 2.3444 | Pass | 10846.5 | 3.036 | 0.027991 | Pass
blank plasma batch 2 |  | 14.996 | 2.2556 | Pass |  | 4.123 | 0.038012 | Pass
blank plasma batch 3 |  | 11.085 | 1.6674 | Pass |  | 0.238 | 0.0021943 | Pass
blank plasma batch 4 |  | 3.229 | 0.48569 | Pass |  |  | 0
blank plasma batch 5 |  | 2.765 | 0.4159 | Pass |  | 2.067 | 0.019057 | Pass
zero volunteer plasma |  | 4.376 | 0.65822 | Pass |  | 2.322 | 0.021408 | Pass
haemolyzed plasma |  |  | 0 | Pass |  | 0.186 | 0.0017148 | Pass
Lipemic plasma |  | 0.903 | 0.13582 | Pass |  | 2.719 | 0.025068 | Pass
concomitant medication |  | 3.495 | 0.5257 | Pass |  | 0.526 | 0.0048495 | Pass
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F09/V1.0 Eff. Date: 22/4/2018

## carry over
Carry over
active ingredient/s | Meropeneme
Calibration no. | 3
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A
Batch name | 1st day validation ( Cal -2 )
Date | 2019-07-05
SN | Area of drug at LLOQ | Area of interfering peak of drug | % | Pass/Fail | Area of Internal standard | Area of interfering peak of Internal | % | Pass/Fail
1 | 664.827 | 18.347 | 2.7597 | Pass | 10846.5 | 3.684 | 0.033965 | Pass
2 |  | 16.721 | 2.5151 | Pass |  | 1.43 | 0.013184 | Pass
3 |  | 26.917 | 4.0487 | Pass |  | 2.733 | 0.025197 | Pass
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC Specialist
Date
SOP/LB/07/F10/V1.0 Eff. Date: 22/4/2018
Acceptance criteria:-
Carry over in the blank sample following the high concentration standard should not be greater than 20% of the lower limit of quantification and 5% for the internal standard.

## Freeze & Thaw Stability 1st
Freeze-thaw Stability Test -1st cycles
active ingredient/s | Meporeneme
Calibration no. | 3 |  | 6 |  |  | 3 |  | 6
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01 |  |  | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A |  | N/A |  |  | N/A |  | N/A
Batch name | 1st day validation (Cal -3 ) |  | 2nd day validation ( Cal -6 ) |  |  | 1st day validation (Cal -3 ) |  | 2nd day validation ( Cal -6 )
Date | 2019-07-05 |  | 2019-07-05 |  |  | 2019-07-05 |  | 2019-07-05
Level | Low QC |  |  |  |  | High QC
Concentration | 1.5 |  |  |  |  | 40
ng/ml | fresh | A% | After 1st cycle | A% | Pass/Fail | fresh | A% | After 1st cycle | A% | Pass/Fail
| 1.312 | 87.45 | 1.29 | 86.005 | Pass | 43.754 | 109.386 | 45.049 | 112.624 | Pass
| [1.243] | 82.881 | 1.32 | 88.022 | Pass | [48.106] | 120.264 | 45.582 | 113.956 | Pass
| [1.221] | 81.37 | 1.287 | 85.824 | Pass | 45.578 | 113.944 | [46.733] | 116.832 | Fail
| 1.301 | 86.761 | [1.177] | 78.454 | Fail | [46.629] | 116.572 | 45.099 | 112.748 | Pass
| 1.322 | 88.125 | 1.28 | 85.351 | Pass | 43.261 | 108.152 | [50.283] | 125.707 | Fail
| 1.308 | 87.214 | 1.333 | 88.883 | Pass | 45.064 | 112.66 | 45.384 | 113.461 | Pass
Mean | 1.3108 |  | 1.302 |  |  | 44.4142 |  | 45.2785
Accuracy% | 87.3833 |  | 86.8 |  |  | 111.0356 |  | 113.1962
CV% | 0.6691 |  | 1.7757 |  |  | 2.4466 |  | 0.55307
Stability(in%) | 99.3324 |  |  |  |  | 101.9459
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F11/V1.0 Eff. Date: 22/4/2018

## Freeze & Thaw Stability 2nd cyc
Freeze-thaw Stability Test -1st cycles
active ingredient/s | Meporeneme
Calibration no. | 3 |  | 8 |  |  | 3 |  | 8
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01 |  |  | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A |  | N/A |  |  | N/A |  | N/A
Batch name | 1st day validation (Cal -3 ) |  | 3rd day validation ( Cal -8 ) |  |  | 1st day validation (Cal -3 ) |  | 3rd day validation ( Cal -8 )
Date | 2019-07-05 |  | 2019-07-08 |  |  | 2019-07-05 |  | 2019-07-08
Level | Low QC |  |  |  |  | High QC
Concentration | 1.5 |  |  |  |  | 40
ng/ml | fresh | A% | After 2nd cycle | A% | Pass/Fail | fresh | A% | After 2nd cycle | A% | Pass/Fail
| 1.312 | 87.45 | 1.339 | 89.265 | Pass | 43.754 | 109.386 | 41.387 | 103.467 | Pass
| [1.243] | 82.881 | 1.28 | 85.33 | Pass | [48.106] | 120.264 | 41.419 | 103.548 | Pass
| [1.221] | 81.37 | 1.299 | 86.572 | Pass | 45.578 | 113.944 | 42.981 | 107.451 | Pass
| 1.301 | 86.761 | 1.303 | 86.86 | Pass | [46.629] | 116.572 | 42.217 | 105.542 | Pass
| 1.322 | 88.125 | 1.36 | 90.698 | Pass | 43.261 | 108.152 | 41.173 | 102.933 | Pass
| 1.308 | 87.214 | 1.296 | 86.426 | Pass | 45.064 | 112.66 | 42.765 | 106.912 | Pass
Mean | 1.3108 |  | 1.3128 |  |  | 44.4142 |  | 41.9903
Accuracy% | 87.3833 |  | 87.5222 |  |  | 111.0356 |  | 104.9758
CV% | 0.6691 |  | 2.3 |  |  | 2.4466 |  | 1.842
Stability(in%) | 100.1589 |  |  |  |  | 94.5425
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F11/V1.0 Eff. Date: 22/4/2018

## Freeze & Thaw Stability 3rd cyd
Freeze-thaw Stability Test -1st cycles
active ingredient/s | Meporeneme
Calibration no. | 3 |  | 9 |  |  | 3 |  | 9
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01 |  |  | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A |  | N/A |  |  | N/A |  | N/A
Batch name | 1st day validation (Cal -3 ) |  | 4th day validation ( Cal -9 ) |  |  | 1st day validation (Cal -3 ) |  | 4th day validation ( Cal -9 )
Date | 2019-07-05 |  | 2019-07-09 |  |  | 2019-07-05 |  | 2019-07-09
Level | Low QC |  |  |  |  | High QC
Concentration | 1.5 |  |  |  |  | 40
ng/ml | fresh | A% | After 3rd cycle | A% | Pass/Fail | fresh | A% | After 3rd cycle | A% | Pass/Fail
| 1.312 | 87.45 | 1.386 | 92.401 | Pass | 43.754 | 109.386 | 45.114 | 112.784 | Pass
| [1.243] | 82.881 | 1.356 | 90.415 | Pass | [48.106] | [120.264] | [46.548] | 116.37 | Fail
| [1.221] | 81.37 | 1.357 | 90.485 | Pass | 45.578 | 113.944 | 44.562 | 111.405 | Pass
| 1.301 | 86.761 | 1.315 | 87.64 | Pass | [46.629] | [116.572] | 45.459 | 113.647 | Pass
| 1.322 | 88.125 | [1.268] | 84.526 | Fail | 43.261 | 108.152 | 44.019 | 110.047 | Pass
| 1.308 | 87.214 | [1.265] | 84.343 | Fail | 45.064 | 112.66 | 45.237 | 113.094 | Pass
Mean | 1.3108 |  | 1.3535 |  |  | 44.4142 |  | 44.8782
Accuracy% | 87.3833 |  | 90.2333 |  |  | 111.0356 |  | 112.1955
CV% | 0.6691 |  | 2.157 |  |  | 2.4466 |  | 1.2992
Stability(in%) | 103.2615 |  |  |  |  | 101.0446
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F11/V1.0 Eff. Date: 22/4/2018

## Short term stability
xh-room temperature-short-term stability test
active ingredient/s | Meropeneme
Calibration no. | 3 |  | 5 |  |  | 3 |  | 5
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL) |  |  | Microlit - Zi\DI\MI\02 (10 mL) |  | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01 |  |  | Eppendorf - Zi/CE/EP/01 |  | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A |  | N/A |  |  | N/A |  | N/A
Batch name | 1st day validation (Cal -3 ) |  | 2nd day validation ( Cal -5 ) |  |  | 1st day validation (Cal -3 ) |  | 2nd day validation ( Cal -5 )
Date of preparation | 2019-07-05 |  | 2019-07-05 |  |  | 2019-07-05 |  | 2019-07-05
Date of Injection | 2019-07-05 |  | 2019-07-05 |  |  | 2019-07-05 |  | 2019-07-05
Level | Low QC |  |  |  |  | High QC
Concentration | 1.5 |  |  |  |  | 40
ng/ml | fresh | A% | After x hrs | A% | Pass/Fail | fresh | A% | After x hrs | A% | Pass/Fail
| 1.312 | 87.45 | 1.283 | 85.56 | Pass | 43.754 | 109.386 | 43.262 | 108.154 | Pass
| [1.243] | 82.881 | 1.281 | 85.384 | Pass | [48.106] | 120.264 | 44.365 | 110.914 | Pass
| [1.221] | 81.37 | 1.303 | 86.887 | Pass | 45.578 | 113.944 | 45.015 | 112.538 | Pass
| 1.301 | 86.761 | 1.311 | 87.41 | Pass | [46.629] | 116.572 | 44.357 | 110.892 | Pass
| 1.322 | 88.125 | [1.271] | 84.753 | Fail | 43.261 | 108.152 | 43.404 | 108.511 | Pass
| 1.308 | 87.214 | [1.229] | 81.909 | Fail | 45.064 | 112.66 | 43.623 | 109.058 | Pass
Mean | 1.3108 |  | 1.2945 |  |  | 44.4142 |  | 44.0043
Accuracy% | 87.3833 |  | 86.3 |  |  | 111.0356 |  | 110.0108
CV% | 0.6691 |  | 1.1449 |  |  | 2.4466 |  | 1.5522
Stability(in%) | 98.7603 |  |  |  |  | 99.0771
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F12/V1.0 Eff. Date: 22/4/2018

## Dilution integrity
Dilution integrity validation report
active ingredient/s |  |  | Meropeneme
Calibration no. | 2
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A
Batch name | 1st day validation ( Cal -2 )
Date | 43651
Balance used & code | Ohous - Zi/BA/OH/01
Level | X2 |  |  | X4
Concentartion | 40 |  |  | 40
ng/ml | Conc (ng/ml) | A% | Pass/Fail | Conc(ng/ml) | A% | Pass/Fail
| [46.371] | 115.927 | Fail | 43.78 | 109.451 | Pass
| 45.817 | 114.543 | Pass | 41.579 | 103.947 | Pass
| 45.083 | 112.707 | Pass | 39.89 | 99.726 | Pass
| 42.515 | 106.288 | Pass | 39.779 | 99.449 | Pass
| 44.786 | 111.965 | Pass | 40.331 | 100.828 | Pass
| 42.569 | 106.421 | Pass | 39.438 | 98.595 | Pass
Mean | 44.154 |  |  | 40.7995
Accuracy% | 110.385 |  |  | 101.9988
CV% | 3.4397 |  |  | 4.0174
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F17/V1.0 Eff. Date: 22/4/2018

## recovery
Analyte Recovery
active ingredient/s | Meropeneme
Date | 2019-07-08 |  | Date | 2019-07-08
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  | Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL) |  | Dispenser used | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01 |  | Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A |  | Concentrator code | N/A
Batch name | Recovery |  | Batch name | Recovery
dilution | 1 | 1 | dilution | 1 | 1
Analyte Recovery Test Report |  |  | Internal standard Recovery Test Report
| Low QC |  |  | Low QC
| Unextracted | Extracted |  | Unextracted | Extracted
| 5732.62 | 5148.67 |  | 42371.9 | 37325.7
| 6211 | 4783.29 |  | 43692.1 | 37600.9
| 6364.77 | 4807.94 |  | 44299 | 37866.5
| 6500.08 | 4748.34 |  | 44340.6 | 37350
| 6425.93 | 4628.82 |  | 44633.6 | 36728.8
| 6172.78 | 4932.94 |  | 44175.5 | 37560.6
Mean | 6234.53 | 4841.6667 | Mean | 43918.7833 | 37405.4167
CV (%) | 4.4241 | 3.7079 | CV (%) | 1.862 | 1.03
Absolute Recovery |  | 77.6589 | Absolute Recovery |  | 85.1695
| Medium QC |  |  | Medium QC
| Unextracted | Extracted |  | Unextracted | Extracted
| 114018 | 86286.6 |  | 44942.4 | 34658.3
| 109834 | 86838.3 |  | 43135.6 | 33968.9
| 110594 | 91212.5 |  | 43403.1 | 34844.9
| 108503 | 88611.8 |  | 43856.8 | 34475.2
| 110511 | 89468.5 |  | 43666.9 | 34203.3
| 109755 | 88720.7 |  | 43745 | 34490.2
Mean | 110535.8333 | 88523.0667 | Mean | 43791.6333 | 34440.1333
CV (%) | 1.6859 | 2.0217 | CV (%) | 1.4178 | 0.91204
Absolute Recovery |  | 80.0854 | Absolute Recovery |  | 78.6455
| High QC |  |  | High QC
| Unextracted | Extracted |  | Unextracted | Extracted
| 274808 | 212943 |  | 39067 | 31864.9
| 269229 | 211654 |  | 38083.4 | 32108
| 267493 | 201177 |  | 38958 | 31555.7
| 272543 | 217014 |  | 38791.3 | 32469.4
| 262092 | 214001 |  | 38506.8 | 32346.2
| 261731 | 213429 |  | 38769.8 | 32138.3
Mean | 267982.6667 | 211703 | Mean | 38696.05 | 32080.4167
CV (%) | 1.9953 | 2.5769 | CV (%) | 0.91869 | 1.0318
Absolute Recovery |  | 78.9988 | Absolute Recovery |  | 82.9036
Comments:
Approval signature
Analyst
Date
Lab. Supervisor
Date
IPC specialist
Date
SOP/LB/07/F18/V1.0 Eff. Date: 22/4/2018

## Matrix
IS Recovery
active ingredient/s | Meropeneme
Date | 2019-07-09 |  |  |  | Date | 2019-07-09
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL) |  |  |  | Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL) |  |  |  | Dispenser used | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01 |  |  |  | Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A |  |  |  | Concentrator code | N/A
Batch name | Matrix |  |  |  | Batch name | Matrix
Analyte Matrix effect Test Report |  |  |  |  | IS Matrix effect Test Report
| Low QC |  |  |  |  | Low QC
| Unextracted | Extracted | Matrix factor | Normalized factor |  | Unextracted | Extracted | Matrix factor
| 4622.27 | 4620.35 | 99.6727 | 1.0006 |  | 36032.5 | 35468.3 | 99.6088
| 4717.19 | 4398.21 | 94.8806 | 0.95408 |  | 35226.3 | 35410.6 | 99.4467
| 4651.28 | 4266.14 | 92.0315 | 0.931 |  | 35420.2 | 35199.1 | 98.8528
| 4607.47 | 4585.44 | 98.9196 | 0.98843 |  | 36009.6 | 35635.3 | 100.0778
| 4742.18 | 4445.8 | 95.9073 | 0.97045 |  | 36110.8 | 35190.1 | 98.8275
| 4472.73 | 4549.1 | 98.1357 | 0.97104 |  | 34846.2 | 35986 | 101.0627
|  | 4822.37 | 104.0308 | 1.0112 |  |  | 36630.8 | 102.8735
|  | 4585.58 | 98.9227 | 0.97496 |  |  | 36128.5 | 101.4629
Mean | 4635.52 | 4534.1237 | 97.8126 | 0.97523 | Mean | 35607.6 | 35706.0875 | 100.2766
CV (%) | 2.0637 | 3.6727 | 3.6727 | 2.6176 | CV (%) | 1.4623 | 1.4139 | 1.4139
| High QC |  |  |  |  | High QC
| Unextracted | Extracted | Matrix factor | Normalized factor |  | Unextracted | Extracted | Matrix factor
| 213098 | 212118 | 100.242 | 1.0051 |  | 31753.7 | 31477.6 | 99.729
| 209012 | 205592 | 97.1579 | 0.96238 |  | 31761.2 | 31864.8 | 100.9557
| 211229 | 213880 | 101.0746 | 0.9822 |  | 31741.2 | 32480.4 | 102.9061
| 218320 | 185324 | 87.5797 | #VALUE!
| 208225 | 231669 | 109.4813 | 1.1248 |  | 31382.9 | 30723 | 97.3382
| 209752 | 224475 | 106.0816 | 1.0747 |  | 31176.7 | 31155.5 | 98.7085
Mean | 211606 | 212176.3333 | 100.2695 | #VALUE! | Mean | 31563.14 | 31540.26 | 99.9275
CV (%) | 1.7561 | 7.5953 | 7.5953 | #VALUE! | CV (%) | 0.85171 | 2.1322 | 2.1322
Comments:
Approval signature
Analyst
Date
Lab. Supervisor
Date
IPC specialist
Date
SOP/LB/07/F19/V1.0 Eff. Date: 22/4/2018

## system suitability
System suitability
active ingredient/s | Meropenem
Calibration no. | 2
Pipette type and code | Zi\MP\EP\04 (1 mL) - Zi\MP\IS\02 (200 µL)
Dispenser used | Microlit - Zi\DI\MI\02 (10 mL)
Centrifuge code | Eppendorf - Zi/CE/EP/01
Concentrator code | N/A
Batch name | 1st day validation ( Cal -2 )
Date | 2019-07-05
Balance used & code | Ohous - Zi/BA/OH/01
SN# | Drug |  | Internal standard
| Peak Area | Retention time(min) | Peak Area | Retention time(min)
1 | 27172.1 | 0.38 | 26783.8 | 0.36
2 | 27882.8 | 0.38 | 26761.2 | 0.36
3 | 27666.6 | 0.38 | 26516.3 | 0.36
4 | 28293.5 | 0.38 | 26965.8 | 0.35
5 | 28185.8 | 0.38 | 26799.9 | 0.35
6 | 28227.9 | 0.38 | 26697.6 | 0.35
Mean | 27904.7833 | 0.38 | 26754.1 | 0.355
CV% | 1.5443 | 1.6002e-14 | 0.54829 | 1.5429
Comments:
Approval signature
Analyst
Date
Lab Supervisor
Date
IPC specialist
Date
SOP/LB/07/F20/V1.0 Eff. Date: 22/4/2018